## Tier 1
Parameter | Equipment | Units of Measurements | Minimum Detection Limit | Accuracy | Resolution | Range | Calibration | Maintenance Needed | General Cost | Manufacturer | Website
Temperature | Taylor Precision Digital Thermometer | Celsius | -5 | ± 1.0 | 0.01 | 230 to -40 C | NSF certified thermometer | Make sure batteries are working properly. | 15 | Various | http://taylor-enviro.com/kitchen/thermometers/5-commercial-anti-microbial-instant-read-1.html
pH | pH indicator strips |  | 5 | ± 1.0 | 0.5 | 5.0 to 10.0 in 0.5 increments | None | One time use; Do not let strips get wet before testing. | 15 | EMD | http://www.emdmillipore.com/US/en/product/pH-indicator-strips--pH-5.0---10.0,MDA_CHEM-109533
Stream Turbidity | Turbidity Tube | cm / NTU | 60 cm or 6 NTU | ± 1.0 cm | 0.2 cm or ~ 2 NTU | 0 to 60 cm / 6 to >240 NTU | None | Check for cracks and keep clean. | 40 | Water Monitoring Equipment & Supply | http://www.watermonitoringequip.com/pages/stream.html
Lake Turbidity | Secchi Disk | Meters | 0.05 meters | ± 0.05 meters | 0.01 meter | 0.0 to 15.0 meters | None | Check that line is secure and disk is clean | 80 | Water Monitoring Equipment & Supply | http://www.watermonitoringequip.com/pages/lake.html
Dissolved Oxygen | Chemetrics | mg/L or ppm | 1 ppm | ± 1 color standard increment | 1 ppm | 1,2,3,4,5,6,8,10,12 | None | One time use. Keep colorchart out of long periods of direct sun to prevent fading. | 53 | Chemetrics | https://www.chemetrics.com/Oxygen+%28dissolved%29/Visual+Kits/K-7512
E.coli | Coliscan Easygel | cfu/100 mL | 20 cfu/ 100 mL |  | 20 cfu/ 100 mL | 20 to ~2400 cfu/ 100 mL | None | Easygel mediums need to remain frozen until testing. Shelf life of one year. Single use test requires sterile whirlpak bag for collection and pipettes for transfer. 8 hour holding period for samples | $27 per set of ten tests | Micrology Laboratories | http://www.micrologylabs.com/products/2/Coliscan-Easygel

## Tier 2
Parameter | Equipment | Units of Measurements | Minimum Detection Limit | Accuracy | Resolution | Range | Calibration | Maintenance Needed | General Cost | Manufacturer | Website
Air Temperature | Enviro Safe Thermoeter | Celsius | -5 | ± 1.0 | 0.5 | 50 to - 5 C | Factory calibration to NIST when purchased. | Check for no air bubbles or cracks | 11 | Various | http://forestrysuppliers.com/product_pages/Products.asp?mi=53561&itemnum=77141
Water Temperature | Oakton PCS Testr 35 | Celsius | 0 | ± 0.5 | 0.1 | 0 to 50 C | 1 point calibration | Store in case after use. Ensure proper batteries. | 160 | Oakton | http://www.4oakton.com/proddetail.asp?parent=2&prod=352&seq=2&Totrec=13
pH | Oakton PCS Testr 35 |  | 0.1 | ± 0.1 | 0.01 | 0.0 to 14.0 | 3 point calibration with pH standards | Probe must be calibrated within a 24 hour period before monitoring. Store with moisture in cap to prevent pH sensor from drying out. Replace sensor when it will not correctly calibrate | 160 | Oakton | http://www.4oakton.com/proddetail.asp?parent=2&prod=352&seq=2&Totrec=13
Conductivity | Oakton PCS Testr 35 | 0.0-2000 µS/cm; 2.01- 20.00 mS/cm | 10 µS/cm | ±1% full-scale | 10 µS/cm | 0.0-2000 µS/cm; 2.01- 20.00 mS/cm | 1 point calibration | Probe must be calibrated within a 24 hour period before monitoring. | 160 | Oakton | http://www.4oakton.com/proddetail.asp?parent=2&prod=352&seq=2&Totrec=13
Total Dissolved Solids | Oakton PCS Testr 35 | 0 to 1000 ppm / 1.00 to 10.0 ppt | 1 ppm | ±1% full-scale | 1 ppm | 0 to 1000 ppm / 1.00 to 10.0 ppt | 1 point calibration | Salinity is derived from a relationship with temperature and conductivity. Probe must be calibrated for conductivity within a 24 hour period before monitoring. | 160 | Oakton | http://www.4oakton.com/proddetail.asp?parent=2&prod=352&seq=2&Totrec=13
Salintity | Oakton PCS Testr 35 | ppm | 1 ppm | ±1% full-scale | 0.1 below 100 ppm; 1 ppm above | 0 to 999 ppm | 1 point calibration | TDS is derived from a relationship with temperature and conductivity. Probe must be calibrated for conductivity within a 24 hour period before monitoring. | 160 | Oakton | http://www.4oakton.com/proddetail.asp?parent=2&prod=352&seq=2&Totrec=13
Stream Turbidity | Turbidity Tube | cm / NTU | 60 cm or 6 NTU | ± 1.0 cm | 0.2 cm or ~ 2 NTU | 0 to 60 cm / 6 to >240 NTU | None | Check for cracks and keep clean. | 40 | Water Monitoring Equipment & Supply | http://www.watermonitoringequip.com/pages/stream.html
Lake Turbidity | Secchi Disk | Meters | 0.05 meters | ± 0.05 meters | 0.01 meter | 0.0 to 15.0 meters | None | Check that line is secure and disk is clean | 80 | Water Monitoring Equipment & Supply | http://www.watermonitoringequip.com/pages/lake.html
Dissolved Oxygen | LaMotte Dissolved Oxygen Winkler Titration | mg/L or ppm | 0.2 ppm | ± 0.5 ppm | 0.2 ppm | 0 to 14 | None | One time use of reagents. Make sure all reagents are not expired. | 55 | LaMotte | http://www.lamotte.com/en/water-wastewater/individual-test-kits/5860-01.html
E.coli | Coliscan Easygel | cfu/100 mL | 20 cfu/ 100 mL |  | 20 cfu/ 100 mL | 20 to ~2400 cfu/ 100 mL | None | Easygel mediums need to remain frozen until testing. Shelf life of one year. Single use test requires sterile whirlpak bag for collection and pipettes for transfer. 8 hour holding period for samples | $27 per set of ten tests | Micrology Laboratories | http://www.micrologylabs.com/products/2/Coliscan-Easygel
E.coli | IDEXX | cfu/100 mL | 1 cfu / 100 mL | 95% of time the test accuratly identifies ( based on reseach studies provied by IDEXX) | 1 cfu per 100 mL | <1 to >2419 cfu per 100 mL | None | Reagents are single time use; make sure none are expired. 8 hour holding period for samples | ~$ 9.25 per test | IDEXX | https://www.idexx.com/water/products/quanti-tray.html